## Slide 1
NATIONAL UNIVERSITY OF SCIENCE AND TECHNOLOGY POLITEHNICA BUCHAREST
Developing a System for Physiological Signal Acquisition and Analysis for Sleep Quality Assessment
DIPLOMA PROJECT
GRADUATE
Duta Alexandru-Vlad
SCIENTIFIC COORDINATOR
Ph.D. Ovidiu Grigore
Faculty of Electronics, Telecommunications and Information Technology · 2026

## Slide 2
00
Presentation Structure
1
Introduction to Sleep
Sleep stages, cycles, and why monitoring matters
2
Motivation
PSG vs commercial wearables vs this project
3
System Architecture & Hardware
ESP32-C3, sensors, wearable integration
4
Data Acquisition & Dashboard
Firmware, server, SQLite storage, live monitoring
5
Signal Processing & Features
Filtering, 21 features per 30-second window
6
Sleep Stage Classification
K-Means clustering, physiological rules, KNN
7
Results & Validation
10 sessions, hypnograms, quality scores
8
Conclusions & Future Work
Achievements, limitations, next steps
Overview
02 / 17

## Slide 3
01
Introduction to Sleep
Sleep occupies roughly one third of human life and is essential for physical recovery, memory consolidation, and emotional regulation. The AASM classifies sleep into four stages based on brain, muscle, and eye activity.
N1
5–10%
Light Transition
Theta waves replace alpha. Transition from wakefulness. Easy to wake.
N2
45–55%
Light Sleep
Sleep spindles and K-complexes. Heart rate falls. Memory consolidation begins.
N3
20–25%
Deep / SWS
Delta waves dominate. Lowest heart rate. Peak growth hormone and immune activity.
REM
20–25%
Rapid Eye Movement
Dreaming occurs. Muscle tone suppressed. Emotional and procedural memory processing.
A typical night contains 4–6 cycles of ~90 minutes each. N3 dominates early cycles; REM periods lengthen toward morning.
Introduction to Sleep
03 / 17
→ System: Light Sleep
Not distinguished from N2 without EEG
→ System: Light Sleep
Not distinguished from N2 without EEG
→ System: Deep Sleep
BPM≤P40 & motion≤P50, capped 25%
→ System: REM Sleep
BPM≥P65 & std≥P55, capped 30%
This system outputs 4 classes — Wake/Movement, Light Sleep (N1+N2 combined), Deep Sleep, REM — since wrist sensors without EEG cannot separate N1 from N2.

## Slide 4
€600 – 2 000
per night
Clinical PSG
Gold standard but expensive, lab-bound, requires trained staff
65 – 75%
accuracy
Commercial Wearables
Affordable but black-box algorithms, no raw data access
<€100
OPEN-SOURCE
This Project
Full raw data access, 5 sensors + ML pipeline
04 / 17
Goal: build an open, low-cost wearable sleep monitoring system with full raw data access
Motivation
Motivation
02

## Slide 5
03
System Architecture
System Architecture & Hardware
05 / 17

## Slide 6
03
Hardware Design
ESP32-C3 Super Mini
RISC-V 160 MHz, WiFi, single-core
MCU
MAX30102
PPG + SpO₂ (red 660nm, IR 880nm)
Optical
MPU6050
6-axis IMU (accel + gyro)
Motion
BMP280
Skin temperature (±0.5°C)
Temperature
ADS1015 + Flex
Respiration via chest band
Respiration
MAX4466
Microphone for snoring detection
Acoustic
TP4056 + LiPo 2500 mAh · >17h runtime
System Architecture & Hardware
06 / 17

## Slide 7
03
Wearable Integration
Wrist Unit
PPG, IMU, temperature, microphone sensors during recording
Elastic Chest Band
Flex sensor for respiration monitoring
System Architecture & Hardware
07 / 17

## Slide 8
1
Sensor Reading
PPG (IR+Red)
12.5 Hz
IMU (Acc+Gyro)
10 Hz
Flex (Resp.)
4 Hz
Temperature
1 Hz
Microphone
200 Hz
→
2
JSON Assembly
Single combined HTTP POST every 2s (single-core constraint)
{ "t":"23:35:01",
"ir":54231,
"red":48120,
"ax":0.01,
"ay":-0.99,"az":0.03,
"fr":612,"fd":-8,
"tmp":35.3,
"ppg_ir":"54231,...",
"mic_s":"1066,..." }
→
3
SQLite Storage
Raspberry Pi stores all raw data in structured tables
raw_data
timestamp, JSON
ppg_data
IR, Red, BPM
flex_data
ADC, resp. rate
mic_data
RMS amplitude
→
4
Analysis Pipeline
21 features/window. K-Means+KNN. Quality score.
Step A
Signal filtering
(sensor-specific)
Step B
Feature extraction
(21 features)
Step C
K-Means clustering
(k=4)
Step D
KNN + smoothing
+ scoring
~227
samples/sec across
all sensors
2s
POST interval
single-core ESP32
6–10h
recording duration
per session
WiFi
HTTP POST :5000
to Raspberry Pi
04
Data Acquisition
Key design decision: single combined POST per cycle avoids sample loss on single-core ESP32-C3
Data Acquisition & Dashboard
08 / 17

## Slide 9
04
Live Dashboard
Real-time signal monitoring on port 8080 during active recording sessions
Data Acquisition & Dashboard
09 / 17

## Slide 10
05
Signal Processing
PPG (IR)
Butterworth 0.5–3.5 Hz, order 3
PPG chart
Gyroscope
Median filter, kernel size 5
Flex Sensor
Moving average, window 3
Temperature
Butterworth lowpass 0.1 Hz, order 3
Signal Processing & Features
10 / 17

## Slide 11
05
Feature Extraction
21 features per 30-second window · StandardScaler · weighted by physiological relevance
PPG / Cardiac
×1.8
bpm_mean
×1.5
bpm_std
×1.0
bpm_min
×1.0
bpm_max
×1.6
hrv_simple
×1.2
spo2_est
Motion
×2.5
motion_level
×2.2
acc_jerk
×2.0
gyro_energy
×0.5
motion_variation
×1.0
dominant_pos
Respiration
×1.8
resp_rate
×1.4
resp_variability
×1.2
resp_amplitude
×1.0
apnea_flag
Temperature
×0.8
temp_mean
×0.5
temp_variation
Acoustic
score only
snore_events_count
score only
snore_intensity
Preprocessing Pipeline
Raw features → StandardScaler (zero mean, unit variance) → Multiply by weights → K-Means input.
contact_valid_ratio excluded (data leakage risk)
Signal Processing & Features
11 / 17

## Slide 12
06
Sleep Stage Classification
1
K-Means (k=4)
Unsupervised clustering into 4 groups.
Silhouette: k=2 best (0.797)
k=4 used for clinical relevance (0.290)
2
Physiological Labeling
Wake: motion ≥ P95
REM: bpm ≥ P65 AND bpm_std ≥ P55
Deep: bpm ≤ P40 AND motion ≤ P50
Light: all remaining windows
3
KNN (k=5)
Trained on K-Means labels. Distance-weighted Euclidean in 21D. Outputs per-class probabilities.
4
Smoothing
3-epoch majority vote (90s)
Deep cap: 25%, REM cap: 30%
Sleep Stage Classification
12 / 17

## Slide 13
07
Results: Hypnogram Analysis
10 overnight sessions · 8 subjects · 375–602 min per session
Results & Validation
13 / 17

## Slide 14
07
Sleep Quality Score
0–100 scale from 6 components · Duration penalty: up to −15 pts for sessions <6 hours
/30
Architecture
Deep ≥ 25%, REM ≥ 25%, Wake < 15%
/20
PPG Quality
BPM 45–70, SpO₂ ≥ 95%
/20
Respiration
RR 10–18, apnea < 2% of windows
/15
Movement
Median motion < 0.05g
/10
HRV
RMSSD 30–100 ms
/5
Snoring
0 events = full
85–100 Excellent
70–84 Good
55–69 Average
40–54 Poor
Results & Validation
14 / 17

## Slide 15
07
Subjective vs Objective Comparison
Results & Validation
15 / 17

## Slide 16
08
Key Findings & Limitations
Key Findings
→
Motion features: most reliable for sleep staging
→
HRV (RMSSD) + respiratory rate: strongest stage differentiation
→
S12 (OSA subject): 52 apnea + 257 snoring events correctly flagged
→
Subjects unaware of apnea events detected by sensors
→
Scores: 62.8 – 80.4 across valid sessions
Limitations
—
No ground truth (PSG) for epoch-level validation
—
Silhouette score k=4: 0.290 (overlap between stages)
—
Single-core ESP32-C3 limits PPG to ~12.5 Hz effective
—
Sensor displacement during sleep affects signal quality
—
Small sample: 10 sessions, 8 subjects
Key Findings & Limitations
16 / 17
Personal contributions:
Complete hardware assembly from off-the-shelf components (<100 EUR)
ESP32-C3 firmware with multi-sensor acquisition and boot calibration
Flask/SQLite backend + real-time monitoring dashboard
Full Python pipeline: signal processing, 21-feature extraction, K-Means + KNN classification, sleep quality scoring
10 overnight recording sessions with 8 subjects + questionnaire validation
→
→
→
→
→

## Slide 17
08
Conclusions & Future Work
The system demonstrates that open-source, low-cost sleep monitoring with full raw data access is achievable with current off-the-shelf components.
FUTURE WORK
01
Dual-core MCU
Continuous 25 Hz PPG for frequency-domain HRV
02
Rigid Housing
Custom 3D-printed enclosure for stable sensor contact
03
Better Respiratory Sensor
Inductive/impedance-based instead of flex sensor
04
Cross-session Training
KNN evaluated on held-out data for real accuracy
05
PSG Validation
Simultaneous recording for ground truth benchmarking
Thank you!
Conclusions & Future Work
17 / 17

## Slide 18
References
[3] R. B. Berry et al., "The AASM Manual for the Scoring of Sleep and Associated Events, Version 2.6," American Academy of Sleep Medicine, Darien, IL, 2020.
[1] R. Stickgold, "Sleep-dependent memory consolidation," Nature, vol. 437, pp. 1272-1278, 2005.
[11] M. de Zambotti, N. Cellini, A. Goldstone, I. M. Colrain, and F. C. Baker, "Wearable sleep technology in clinical and research settings," Med. Sci. Sports Exerc., vol. 51, pp. 1538-1557, 2019.
[14] M. Marino et al., "Measuring sleep: accuracy, sensitivity, and specificity of wrist actigraphy compared to polysomnography," Sleep, vol. 36, pp. 1747-1755, 2013.
[17] J. Allen, "Photoplethysmography and its application in clinical physiological measurement," Physiol. Meas., vol. 28, pp. R1-R39, 2007.
[19] Task Force of the European Society of Cardiology and the North American Society of Pacing and Electrophysiology, "Heart rate variability: standards of measurement, physiological interpretation and clinical use," Circulation, vol. 93, 1996.